## Slide 1: Endovascular Stenting of a Left Ventricular Assist Device (LVAD) Outflow Graft Stenosis
Will Johnson, MD
Neil Halin, DO
Tufts Medical Center
Cardiovascular and Interventional Radiology
NESIR Meeting 2/13/2017

## Slide 2: Patient Case
47 Y M hx of NICM, LVEF 10-15%, EtOH abuse
  Heartware LVAD as bridge to transplant 3/13/2014
  Heartware LVAD Exchange due to thrombus 2/18/2015
  Persistent low flows on LVAD
    Currently in 2-3 L/min range; previously had been 3.5 to 4.5 L/min
  Admitted to Tufts MC for further workup
  CTA Chest obtained

## Slide 3: Heartware LVAD

## Slide 8: LVAD Outflow Graft Stenosis
Multidisciplinary Meeting
-Cardiothoracic Surgery
-Interventional Radiology
-Heart Failure
LVAD exchange with CT surgery
  Thoracotomy/open surgery
  High morbidity
  Second LVAD exchange in < 2 years
Endovascular Approach
  Lower morbidity
  Familiarity with approach
  No thrombus seen on CT

## Slide 14: LVAD Outflow Graft Stenosis
Balloon Expandable Bare Metal Stent
  No thrombus
  Appropriate for kink
Covered Stent
  Thrombus
  Pseudoaneurysm
  Hematoma

## Slide 15: LVAD Outflow Graft Stenosis
Balloon Expandable Stents in areas of stenosis
  10 mm x 37 mm Boston Scientific (proximal stenosis)
  10 mm x 37 mm Visi-Pro Balloon (distal stenosis)
LVAD motor RPM decreased during ballooning
LVAD RPM’s returned to normal following stent placement

## Slide 21: Post Procedural
LVAD flow
  Improved to 4.9 L/min from 1.9 L/min prior to procedure
Pulsatility
  BP waveform demonstrated pulsatility prior to procedure indicating LV participation
  Monophasic, suggests offloading of left ventricle by LVAD
Discharged POD 4
  LVAD flows 5-6 L/min by time of DC

## Slide 22: Post Procedure
Outpatient F/U 1 month post op
  Improved energy
  Mild dyspnea on exertion, otherwise minimal heart failure symptoms
  One dose of torsemide due to dietary indiscretion
  Improved flow (one low flow alarm)
  No hospitalizations
  UNOS 1A for heart transplant

## Slide 23: Literature Review

## Slide 24: Literature Review

## Slide 25: Acknowledgements
Neil Halin, DO
Tufts Medical Center, Cardiovascular and Interventional Radiology
Tufts Medical Center, Department of Radiology

## Slide 26: References
Pham, D.T., Kapur, N.K., Dermody, M, Halin, N. (2015). Stenting of an outflow graft obstruction after implantation of a continuous-flow, axial-flow left ventricular assist device. The Journal of Thoracic and Cardiovascular Surgery, 150, e11-12
Ganapathi, A., Andersen, N., Prastein, D., Hashmi, Z., Rogers, J., Milano, C., McCann, R., Hughes, G.C. (2013) Endovascular Stent Grafting of a Left Ventricular Assist Device Outflow Graft Pseudoaneurysm. Circ Heart Fail, 6, e16-e18
Hanke, J., ElSherbini, A., Rojas, S., Avsar, M., Shrestha, M., Schmitto, J. (2016) Letters to the Editor: Aortic Outflow Graft Stenting in Patient with Left Ventricular Assist Device Outflow Graft Thrombosis. Artificial Organs, 40(4): 414-418
Ahmad, F., Sauer, A., Ricciardi, M. (2016) Endovascular Repair of Ventricular Assist Device Outflow Cannula Stenosis. Catheterization and Cardiovascular Interventions. Published online 10/12/2016
El Sabbagh, A. Al-Hijji, M. Gulati, R., Charanjit, S., Pollak, P., Behfar, Atta. Percutaneous Stenting of a Left Ventricular Assist Device Outflow Kink. (2016) Cardiovascular Interventions. 9(24): e229-e231
Abraham, J., Remick, J., Caulfield, T., Puhlman, M., Evenson, K., Ott, G., Kirker, E. (2014) Left Ventricular Assist Device Outflow Cannula Obstruction Treated with Percutaneous Endovascular Stenting. Circ Heart Fail, 8, 229-230